## Slide 1
Luminance value NG of Porsche CID J1 Dual
Contents
1. Define
2. Measure
3. Analyze
4. Improve
5. Control
6 Sigma GB Assignment
IVI LD Nguyen Duc Long
Notes: 1

## Slide 2
Six Sigma Project Register
Team | IVI LD
Leader | Nguyen Dinh Huong TL
Owner | Nguyen Duc Long
Theme | Reduce defect rate of Model Porsche CID J1 Dual Luminance |  |  |  | Period | 2026.Mar.09 ~ 2026.Sep.20 |  |  | Task team name | Inspection Engineering Project
Six Sigma Project |  | Improvement Goal (Stretch Goal) |  |  |  |  |  |  | Idea of breaking the limit | 
 |  | Improvement index |  | Current |  | World Best | World Best | Goal |  | 
Improve Luminance value |  | KPI | Defect rate (ppm) | 226,624ppm |  | N/A | N/A | 0 ppm | Teaching robot and change camera angle for detecting the central of PID Make check sheet to control camera equipment | 
 |  | ProjectY | Luminance (800-2000Cd/m2) | Ppk = 0.45 Z = 1.34 |  | N/A | N/A | Cpk ≥ 1.33 Z = 4.5 |  | 
Management Issue (Biz. Issue) |  | Implement Plan (Fast & Strong) |  |  |  |  |  |  | Task Member | 
- The Luminance issue affects the line UPH and causes loss time. - NG LCDs after retest are returned to the supplier (HEESUNG) for scrap, leading to IF cost, so an action is needed to eliminate this issue and improve factory performance. |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  | Location | 
 |  |  |  |  |  |  |  |  | LG Electronics Viet Nam-Hai Phong | 
Qualitative Performance |  | Quantitative Performance |  |  |  |  |  |  | Neck Point | 
Improving smoothly testing process for sets Reducing cost |  | 1 LCD part =118.72$ The scrap cost of Luminance issue on March: 19*118.72 = 2255.68$  Result: 0 NG / 420 sets after improvement |  |  |  |  |  |  | Because improvement action related to equipment, So need support from PE team | 
 Schedule
Key activity / Output
D : Define issue, define defect rate, collect defective sample and issue information
M : Gage R&R (Type 1, Type 3) to evaluate equipment
A : Use logic tree and 5 Why to find root cause. Test samples with different CANSIM, camera and camera angle
I : Meeting with HQ IE to discuss improvement action. Correct camera angle to 90° and verify the result
C : Control by daily Master sample check; defect rate via GMES
D
M
A
I
C
3/9
4/1
5/1
5/19
6/15
Task category
□ HW □ SW
□ Mecha ■ Current job
Team | Name | Role
IVI LD | Nguyen Duc Long | Task leader
 | Hoang Nhan Quoc | Member
IVI PE | Pham Van Hung | Member
IVI PM | Ta Van Cong | Member
Notes: Define: 3/9~3/31 (about 3 weeks)
Measure: 4/1~4/30 (4 weeks)
Analyze: 5/1~5/18 (about 3 weeks)
Improve: 5/19~6/14 (4 weeks)
Control: 6/15~9/20 (monitoring)
2

## Slide 3
1.1 Y-y breakdown
Big Y
Little y2
Luminance value
Calculate contrast value
Luminance Function
Little y1
Machine defect
Reduce Porsche CID Defect rate
Hardware defect
Software defect
33,376 PPM
260,000 PPM
Decrease defect rate of Porsche CID
Before
After
6 Sigma assignment
Real defect _ Material
Return to supplier
Activity
Little y3
Hardware issues (6.0%, 15,600 PPM) are mainly broken pin connector (check PCBA), foreign material inside the set after assembly, SQA circuit and LCD from supplier; they will be improved by HW members. Software issues (3.5%, 9,100 PPM) will be solved by HQ software members. Machine defect accounts for 90.5% (235,300 PPM), where Luminance value is the overwhelming root issue causing 226,624 PPM (87.2% of all defects), causing severe UPH loss and scrap cost. That is why Luminance value is selected as Project Y for this Six Sigma assignment.
IVI LD IE
HEESUNG vendor
PIC
260,000 PPM (164/631EA Nov.2025~Mar.2026)
90.5% 235,300 PPM
3.5% 9,100 PPM
6.0% 15,600 PPM
Haptic/Touch
Function
1.0% 2,600 PPM
89.5% 232,700 PPM
87.2% 226,624 PPM
2.3% 6,076 PPM
Power Function
Initial/ Booting
Broken pin connector _ check PCBA
Mr. Hai HW
Uniformity Function
Black Uniformity value
Improve software
HQ R&D

## Slide 4
1.2 Pareto
Chart: Pareto chart of Porsche CID defects (Nov.2025~Mar.2026)
| Luminance value | Hardware | Software | Contrast value | Haptic / Touch
Defect rate (ppm) | 226624 | 15600 | 9100 | 6076 | 2600
Cumulative % | 87.2 | 93.2 | 96.7 | 99 | 100
80% line | 80 | 80 | 80 | 80 | 80
80:20 rule
About 80% of the defects come from about 20% of the items (the vital few). Solve the vital few first.
Result
Luminance value = 226,624 ppm = 87.2% of all defects (> 80%): it is the vital few.
The other 4 items together are only 12.8% (Hardware 6.0%, Software 3.5%, Contrast value 2.3%, Haptic/Touch 1.0%).
So Luminance value is selected as Project Y for this 6 Sigma assignment.
Data: 164 NG / 631 EA = 260,000 ppm (Nov.2025~Mar.2026); Luminance value = 143 NG = 226,624 ppm.
Item (80%): Luminance value PIC: IVI LD (this 6 Sigma assignment) MFG: Line IVI 5, FA production line
Other items: Hardware (broken pin connector) → Mr. Hai HW; Software → HQ R&D; Contrast value (real defect) → HEESUNG vendor.

## Slide 5
1.3 Define Project Y
To reduce the Luminance test NG, Project Y is defined as “Luminance value”.
What is Luminance ?
Project Y
Luminance is a photometric quantity used for light sources and displays such as LCDs.
Luminance is the luminous intensity per unit area emitted by a surface in a given direction (unit: cd/m², also called nit)
It quantifies the brightness level of the display.
Project Y: “Luminance value” of the Porsche CID J1 Dual model.
+
Camera
CS2000
PID
CID
J1 Dual PPR RHD set
The sensor inside the detector receives the light from the center of the LCD, calculates the luminance value and returns the measurement result to the program.
Issue: The luminance measured on the PID screen is lower than the average of the compared values, so the luminance test result falls below the spec.
Project Y= Luminance value = f(x1, x2, …, xn) + ε
Measurement method:
The set is connected to the test bench; when the program sends the request, the display shows the highest brightness level.
The J1 display shows a white screen at maximum brightness.
The CS2000 detector moves to the center of the PID LCD, recognizes the center of the display and measures the luminance.
Because the center always has higher brightness than the corner and edges, the detector must be placed exactly at the center.
The result is recorded and sent back to the program.
Luminance formula:
L = d²Φ / (dA · dω · cos θ)
Spec: 800-2000 (cd/m2)
d²Φ stands for the luminous flux in lumens (lm)
θ stands for the angle between the surface normal of the display and the sight line of the detector, in this case θ = 0°
dA stands for the differential area in square meters (m²)
dω stands for the differential solid angle in steradians (sr)

## Slide 6
Equipment information:
Equipment:
- Power supply, Pallet
- CID bench
- CANSIM
- Camera CS2000 (spectroradiometer)
- Resolution: 0.01 Cd/m2
Program: LG Moohantech
Function test Luminance station
Checking brightness level
Detect and measure Luminance value
Test set without CAN mode
Checking CID black / white patterns
Calculate contrast ratio
Control
Improve
Analyze
Define
Measure
1 Preparation of Gage R&R
Camera CS2000 (spectroradiometer)
Luminance station diagram
Model: Porsche CID J1 Dual PPR (PID / CID display set)
Station: Luminance station (camera CS2000), the station where the Luminance NG occurs
How to get sample (sampling)
Type 1: 1 golden sample (reference 910 Cd/m²), measured 30 times.
Type 3: 25 samples, each measured 2 times (the station is automatic, no operator influence).
Guideline: the samples should cover the spec and the dispersion of the process (include NG / OK samples and samples near the spec).
Notes: Output level resolution: 0.01 Cd/m²

## Slide 7
2.1 Type 1 Study
Control
Improve
Analyze
Define
Measure
07 / 28
Luminance station
Conclusion
• P value < 0.001 (≤ 0.05) Reject H0 -> Bias ≠ 0 (FAIL)
• Cg = 0.91 (Std ≥ 1.33) FAIL | Cgk = 0.49 (Std ≥ 1.33) FAIL
• %Var(Repeatability) = 21.99% (> 10%) FAIL
• %Var(Repeatability & Bias) = 40.63% (> 10%) FAIL
No | Luminance(Cd/m2) | No | Luminance(Cd/m2)
1 | 779.90 | 16 | 872.50
2 | 811.25 | 17 | 881.38
3 | 794.00 | 18 | 820.60
4 | 829.24 | 19 | 905.68
5 | 887.44 | 20 | 897.04
6 | 844.39 | 21 | 833.60
7 | 851.46 | 22 | 885.70
8 | 799.79 | 23 | 838.27
9 | 903.47 | 24 | 837.62
10 | 876.02 | 25 | 863.52
11 | 904.05 | 26 | 892.00
12 | 918.50 | 27 | 847.93
13 | 890.05 | 28 | 885.27
14 | 904.20 | 29 | 786.15
15 | 748.30 | 30 | 859.40
Spec: 800.00 ~ 2000.00
Reference value: 910 (golden sample)
Resolution of Instrument: 0.01
I chose the Type 1 gage study to assess the measurement capability of the gage, based on its repeatability and bias.
In this case, one sample (golden sample, 910 Cd/m²) is measured 30 times.
H0: Bias = 0 in measurement system.
H1: Bias ≠ 0 in measurement system.
✖ Equipment has problem. Cannot measure consistently (large repeatability error and significant bias). Need to improve measurement system.
Notes: Step 1: Run chart: all measurements should stay inside Reference ± 10% of the tolerance (910 ± 120 = 790 ~ 1030 Cd/m²).
Step 2: Check the p-value to assess the bias (p-value vs. α = 0.05).
Step 3: Check Cg and Cgk to assess the capability of the measurement system (benchmark: ≥ 1.33).
Cg (potential gage capability): compares 20% of the tolerance with the spread (6 x s) of the repeated measurements, i.e. repeatability only. Criterion: Cg ≥ 1.33.
Cgk (gage capability index): like Cg but also takes the bias into account (distance between the mean of the measurements and the reference value). Criterion: Cgk ≥ 1.33.

## Slide 8
2.2 Type 3 Study
Control
Improve
Analyze
Define
Measure
Spec: 800- 2000 (Cd/m2)
Luminance station
Need to evaluate the repeatability of the measurement system.
Because the Luminance station is fully automated with no operator influence, Type 3 Gage Study is used to evaluate equipment repeatability.
In type 3 study, use 25 samples to verify station, each sample is measured repeat 2 times.
No | Data | No | Data
1 | 796.95 | 13 | 837.73
1 | 786.61 | 14 | 811.78
2 | 794.73 | 14 | 833.46
2 | 796.58 | 15 | 836.42
3 | 838.61 | 15 | 837.85
3 | 839.58 | 16 | 817.70
4 | 810.76 | 16 | 857.90
4 | 830.05 | 17 | 837.17
5 | 815.12 | 17 | 838.29
5 | 835.75 | 18 | 828.88
6 | 823.98 | 18 | 826.94
6 | 834.57 | 19 | 840.14
7 | 827.40 | 19 | 839.72
7 | 827.69 | 20 | 830.75
8 | 839.12 | 20 | 830.78
8 | 835.42 | 21 | 823.24
9 | 839.73 | 21 | 826.32
9 | 842.06 | 22 | 825.35
10 | 834.01 | 22 | 827.79
10 | 834.48 | 23 | 838.87
11 | 834.67 | 23 | 837.29
11 | 836.61 | 24 | 850.33
12 | 786.15 | 24 | 819.42
12 | 780.77 | 25 | 755.89
13 | 827.29 | 25 | 853.94
% Contribution of Gage R&R = 71.56% > 9% (std < 1%) NG
% Study Var = 84.59% > 10%  NG
% Tolerance = 8.33% < 10%  OK
NDC: 1 < 5  NG
Conclusion: %Contribution = 71.56% (> 9% NG), %Study Var = 84.59% (NG), %Tolerance = 8.33% (OK), NDC = 1 (< 5 NG). The repeatability error (SD ≈ 16.7) is large vs. the part variation (SD ≈ 10.5) and Type 1 also failed, so the system is NOT acceptable: improve.
Notes: H0 (ANOVA, part effect): all part means are equal
H1: at least one part mean is different
Follow AIAG guidelines: %Study Var and %Tolerance < 10% = acceptable, 10-30% = conditionally acceptable, > 30% = unacceptable; %Contribution < 1% = acceptable, > 9% = unacceptable. Here %Contribution (71.56%) and %Study Var (84.59%) are far above the limits and NDC = 1 < 5, so the measurement system is NOT acceptable (only %Tolerance = 8.33% is below 10%). The ANOVA p-value (0.077 > 0.05) also shows that the 25 selected parts cannot be distinguished from the repeatability error.
NDC (number of distinct categories): the number of groups of parts the measurement system can reliably distinguish; NDC ≥ 5 is required.

## Slide 9
Control
Improve
Analyze
Define
Measure
2.3 Result of current Cpk
Current Cpk
P-value < 0.005 < 0.05 → NG
(Data do not follow a normal distribution)
Current process capability: Ppk = 0.45 < 1.33 → NG
Lack of process capability
The process mean 826.25 is close to the LSL 800; about 14% of the readings (7 of 50) are below the LSL.
Baseline sample: the 50 readings (25 sets × 2) measured at the station before improvement, see 2.2.
Target of improvement
Improvement goal (practical): measure the luminance correctly, i.e. the reference sample reads 910 Cd/m² (bias → 0).
(Before improvement) Z = 1.34 → (After improvement) Cpk ≥ 1.33, Z = 4.5
Defect rate 226,624 ppm → 0 ppm

## Slide 10
3.1 Root cause of measurement system problem
Logic tree analysis
Luminance at the station
Man
Automatic process
Machine
Pallet
Power supply
Power supply is stable
Working skill
All model used same kind of pallet
Camera equipment (CS2000)
Measure incorrectly
Power supply is not affected
Changing pallet no different
Change new camera
Man is not related
Material
Method
Return LCD and check on vendor  result is OK
Set
LCD is not affected
Sequence test, timing
Sequence is used for all model, but only this model have issue
Sequence is not affected
Other function is OK, only Luminance is NG
Plug in set to pallet not fully
Not relate to set plugging in


: Potential factor
Camera Angle is changed
Change angle of camera
After the Type 1 study, we see that the measurement system of the Luminance station has problems, so a logic tree is used to find the root cause
Control
Improve
Define

CANSIM
CANSIM send weak signal for LCD recognize
Change new CANSIM

Measure
Analyze

## Slide 11
3.1 Root cause of measurement system problem
: Potential factor

2 Samples T
Control
Improve
Measure
Define
Analyze
No | Potential Factors | List checking | Tool to identify
1 | CANSIM send weak signal to LCD recognized | Checking luminance value before /after change to new CANSIM | 2 sample T
2 | Measure incorrectly | Checking luminance value before /after change to new camera | 2 sample T
3 | Camera angle changed | Checking luminance value before /after changing angle of detector | 2 sample T

## Slide 12
3.1 Root cause of measurement system problem
Control
Improve
Measure
Define
Analyze
2 Samples T _ Analysis
Detection Luminance from CANSIM of the Luminance station_Change to the new CANSIM
H0: The changing CANSIM does not affect in the luminance in the Luminance station
H1: Changing the CANSIM affects the luminance in the Luminance station
P-value = 0.695 > 0.05 -> Fail to reject H0 (No significant difference before/after changing CANSIM).
Changing CANSIM has no significant effect on luminance (Factor eliminated).
Both samples fail the normality test (Anderson-Darling P < 0.005); with n = 30 per group the 2-sample t-test (Welch) is robust.

## Slide 13
3.1 Root cause of measurement system problem
Control
Improve
Measure
Define
Analyze
2 Samples T _ Analysis
Detection Luminance from Camera of the Luminance station_Change to the new camera
H0: The changing new camera does not affect in the luminance in the Luminance station
H1: Changing the new camera affects luminance in the Luminance station
Both samples fail the normality test (Anderson-Darling P < 0.005). With n = 30 per group the 2-sample t-test (Welch) is robust: P = 0.950 > 0.05, so changing the camera has no significant effect.
P-value = 0.950 > 0.05 -> Fail to reject H0 (No significant difference before/after changing camera).
Changing camera has no significant effect on luminance (Factor eliminated).

## Slide 14
3.1 Root cause of measurement system problem
: Potential factor

Control
Improve
Measure
Define
Analyze
2 Samples T _ Analysis
Detection Luminance from Camera of the Luminance station_Change to the angle of Camera
H0: The changing angle does not affect in the luminance in the Luminance station
H1: Changing angle affects the luminance in the Luminance station
Changing camera angle significantly affects luminance → Root Cause verified!
P-value < 0.001 (< 0.05) -> Reject H0 (Significant difference before/after angle adjustment).
The before-fix sample (C1) fails the normality test (Anderson-Darling P < 0.005); with n = 30 per group the 2-sample t-test (Welch) is robust. After the angle correction the mean is 48.99 Cd/m² higher (95% CI 32.93 to 65.05).

## Slide 15
3.1 Root cause of measurement system problem
: Potential factor

2 Samples T
Control
Improve
Measure
Define
Analyze

No | Potential Factors | List checking | Tool to identify | P value | Conclusion
1 | CANSIM send weak signal to LCD recognized | Checking luminance value before /after change to new CANSIM | 2 sample T | 0.695 | Not root cause (P > 0.05)
2 | Measure incorrectly | Checking luminance value before /after change to new camera | 2 sample T | 0.950 | Not root cause (P > 0.05)
3 | Camera angle changed | Checking luminance value before /after changing angle of detector | 2 sample T | < 0.001 | Root Cause (Significant effect, P < 0.05)

## Slide 16
Easy Working: ✔ (automatic test, no operator action)
Difficulty Working: –
Impossible Working: –
Working Feeling: Can detect by program
With who: Mr. Chien LQC
Field request: when issues happen, the repair member uses repair station to check if NTF, PE member check the camera status.
Date: Feb-2026
Line: IVI 5
Area: FA production line
Detail Checking: Visual check the cable of LCD with PCBA to check set OK or not
Problem
Field Checking
Field Workability
Field Interview
Improvement
Cause
Progress Status
Progress Monitoring
Owner
Verify
Remark
Schedule
• Issue: Luminance NG ppm (Model: J1 Dual)
• Defect date: Feb-2026 | Station: Luminance
• Q’ty: 92 ea / 383,333 ppm (Total before: 143 ea / 631 = 226,624 ppm)
Jun
Root Cause: 5-Why Analysis
1. Why occur? Luminance value measured under spec (800~2000 Cd/m²).
2. Why under spec? Camera axis not perpendicular (90°) to the LCD surface.
3. Why angle changed? Long production time caused angle drift from 90° standard.
4. Why did angle drift? Dual display (PID/CID) requires camera motion; vibration loosened fixture.
5. Why no prevention? No daily Master sample check sheet to verify angle before production.
[Luminance NG]
1. Teaching robot & adjust angle of CS2000 Camera (Std: 90.0°)
14- Jun
Mr. Hung PE
Improvement
1. Adjust angle
of CS2000
Camera
14- Jun
2. Monitoring
result
Mr. Hung PE
Control
Improve
Define
Measure
Analyze
3.1 Root cause of measurement system problem
2. Add check sheet to measure Master sample and verify Camera angle
3. Inform supplier (Heesung): LCD is not a cause (returned LCD OK at vendor), no process change needed

## Slide 17
Analyze
Control
Measure
Define
Improve
Before measurement XZ: 89.075º ; YZ: 88.806º
After measurement
XZ: 89.889º ; YZ: 89.889º
CID
PID
Z
X
Y
CID
PID
Z
X
Y
4.1 Correct the camera angle (before / after)
After adjusting the detector camera angle to the 90.0° target (measured 89.889°) and re-testing samples, no Luminance NG was observed in the re-test. Correcting CS2000 camera angle is confirmed as the effective solution.
4. Improvement (short term)

## Slide 18
+
Camera
CS2000
PID
CID
J1 Dual PPR RHD
Center
1. Camera position is correct after checking.
 Camera center coincides with the SET center
2. LHD and RHD use the same servo positions (48, 49) but mirrored screens.  If the position is not correct, J1 PPR LHD will be NG.
(Servo positions 48 and 49: see the diagrams below)
Analyze
Control
Measure
Define
Improve
4.1 Correct the camera angle: camera position and servo position
4. Improvement (short term)

## Slide 19
No | IVI 5 LINE |  |  | CS2000 Apply luminance compensation
 | Model |  | BLU Type | 
1 | Bentayga (Bentley) | CID | Public(Normal) | -
2 | Macan LZV | CID | Public(Normal) | -
3 | 992 | CID | Public(Normal) | -
4 | J1 CID | CID | Public(Normal) | -
5 | G3 CID | CID | Public(Normal) | -
6 | E3PA CID | CID | Public(Normal) | -
7* | J1 Dual LHD | PID | Public(Normal) | -
 |  | CID | Public(Normal) | -
 | J1 Dual RHD | PID | Public(Normal) | -
 |  | CID | Public(Normal) | -
8 | Lamborghini | CID | Public(Normal) | -
9 | E3PA Dual LHD | PID | Privacy | -
 |  | CID | Public(Normal) | -
 | E3PA Dual RHD | PID | Privacy | 13%↑
 |  | CID | Public(Normal) | -
10 | G3 Dual LHD | PID | Privacy | -
 |  | CID | Public(Normal) | -
 | G3 Dual RHD | PID | Privacy | 13%↑
 |  | CID | Public(Normal) | -
11 | J1 Dual PPR LHD | PID | Privacy | -
 |  | CID | Public(Normal) | -
 | J1 Dual PPR RHD | PID | Privacy | 13%↑
 |  | CID | Public(Normal) | -
12 | G2 | CID | Public(Normal) | -
13 | E3 | CID | Public(Normal) | -
Privacy View Angle (PID)
Public View Angle (CID)
Public
Privacy

89°
90°

89°
90°
Difference between Privacy and Public: whether or not an LCF (louver) film is applied
for the privacy display function within the BLU module
 Correcting the CS2000 camera angle from the existing 89° to 90° raises the measured luminance of the RHD Privacy (PID) displays by about 13%, whereas Public (CID) models change by less than 0.5%, so only the RHD Privacy models need luminance compensation (13%).
Privacy & Public Brightness changes depending on viewing angle
Analyze
Control
Measure
Define
Improve
4.2 Luminance compensation for Privacy models
4. Improvement (short term)
* Discontinued model (J1 Dual LHD / RHD)

## Slide 20
Analyze
Measure
Define
Improve
Control
4. Improvement (long term)
4.3 Add Master sample check sheet to control the camera angle
A daily check sheet is established to monitor equipment status using a Master sample. Because the J1 Dual model features dual displays requiring camera movement, camera angle status must be verified during pre-production setup to prevent angle deviation.
Before: no daily check of the camera angle, so an angle drift is not detected until NG sets appear.
After: a daily check sheet with the Master sample (reference 910 Cd/m²) at shift start to verify the camera angle.
Actual DID check sheet in production line
DID check sheet format

## Slide 21
4.4 Type 1 Study again after improvement
21 / 28
Spec: 800.00 ~ 2000.00
Reference value: 910 (golden sample)
Luminance station
Conclusion
• P value = 0.926 > 0.05 Fail to reject H0 -> No significant bias (PASS)
• Cg = 2.51 (Std ≥ 1.33) PASS | Cgk = 2.50 (Std ≥ 1.33) PASS
• %Var(Repeatability) = 7.98% (< 10%) PASS
• %Var(Repeatability & Bias) = 8.00% (< 10%) PASS
✔ Equipment is OK. Camera CS2000 of the Luminance station can be used to measure Luminance value.
No | Luminance(Cd/m2) | No | Luminance(Cd/m2)
1 | 913.97 | 16 | 919.07
2 | 910.09 | 17 | 909.74
3 | 934.62 | 18 | 906.54
4 | 926.63 | 19 | 915.48
5 | 901.07 | 20 | 894.75
6 | 921.35 | 21 | 927.84
7 | 931.37 | 22 | 930.79
8 | 872.12 | 23 | 884.59
9 | 909.81 | 24 | 940.66
10 | 903.27 | 25 | 916.63
11 | 898.65 | 26 | 910.92
12 | 898.47 | 27 | 907.23
13 | 895.41 | 28 | 893.01
14 | 907.88 | 29 | 907.52
15 | 919.56 | 30 | 883.01
Analyze
Control
Measure
Define
Improve
Use the Type 1 gage study to assess the measurement capability of the gage, based on its repeatability and bias after improvement.
In this case, one sample (golden sample, 910 Cd/m²) is measured 30 times.
H0: Bias = 0 in measurement system.
H1: Bias ≠ 0 in measurement system.

## Slide 22
Define
Measure
4.4 Type 3 Study again after improvement
Luminance station
% Contribution of Gage R&R = 0.22% < 1%  OK
% Study Var = 4.64% < 10%  OK
% Tolerance = 0.86% < 10%  Can use to judge OK/NG
NDC: 30 > 5  OK
Conclusion : Measurement system of the Luminance station is Good after improving measurement system
After adjusting camera angle to 90.0°, Type 3 Gage Study was repeated for the Luminance station to verify measurement repeatability.
Control
No | Data | No | Data
1 | 946.86 | 13 | 928.45
1 | 949.24 | 14 | 919.77
2 | 871.49 | 14 | 918.29
2 | 871.89 | 15 | 919.34
3 | 846.49 | 15 | 918.22
3 | 847.85 | 16 | 903.63
4 | 901.35 | 16 | 900.43
4 | 898.05 | 17 | 847.94
5 | 951.03 | 17 | 850.22
5 | 950.78 | 18 | 872.57
6 | 890.56 | 18 | 869.91
6 | 892.16 | 19 | 948.43
7 | 855.23 | 19 | 950.59
7 | 851.36 | 20 | 918.35
8 | 889.44 | 20 | 920.88
8 | 886.99 | 21 | 846.80
9 | 882.94 | 21 | 847.89
9 | 882.16 | 22 | 870.83
10 | 933.20 | 22 | 871.74
10 | 937.61 | 23 | 961.99
11 | 861.69 | 23 | 963.88
11 | 859.00 | 24 | 931.05
12 | 854.84 | 24 | 936.46
12 | 856.49 | 25 | 877.87
13 | 926.83 | 25 | 877.12
Analyze
Improve
Notes: Assess the variation of each source and of any additional factor that affects the measurement system.
H0 (ANOVA, part effect): all part means are equal
H1: at least one part mean is different
VarComp %Contribution: almost all of the variation comes from part-to-part, so the measurement system can reliably distinguish between parts
%Study var: compare the measurement system variation to the total variation
Follow AIAG guidelines: if %Study Var and %Tolerance are below 10% (and %Contribution below 1%), the measurement system is acceptable. Here all three are below the limits and NDC ≥ 5, so the measurement system is Good.

## Slide 23
D
I
C
M
A
Control
Define
Cpk = 2.39 (= CPL; USL 2000 is far from the process) (Goal ≥ 1.33; target ≥ 1.67)  Pass
P-value = 0.503 > 0.05 -> Fail to reject H0 -> No evidence against normality (PASS)
(Data can be treated as normally distributed)
4.5 Process Capability review
Luminance station
Analyze
Measure
Improve
To evaluate process capability of the Luminance station:
Normality test (Anderson-Darling): H0: data follow a normal distribution; H1: data do NOT follow a normal distribution.
Before (2.3): Ppk = 0.45 (Z = 1.34) → Standard ≥ 1.33 → NG, lack of process capability
After: Cpk = 2.39, Ppk = 2.30 (Z = 7.18) → Standard ≥ 1.33 → OK, good process capability
Notes: 1. How to select sample for Process Capability? →Select by Rational subgroup
– Lấy theo lot, n con liền nhau của mỗi lot (n<=5), tại các khoảng thời gian cố định. Vd: cách 2 tiếng một lần trong nhiều ngày liên tiếp
Mục đích: để đảm bảo điều kiện
của những sample được lấy là giống hệt nhau, chỉ tồn tại chance cause (nguyên nhân ngẫu nhiên) trong mỗi nhóm.
Tại sao n <= 5? Vì từ con thứ 6 trở đi không đảm bảo điều kiện giống
như con đầu tiên nữa.
How to choose samples for process capability? Take n consecutive units from each lot (n ≤ 5) at fixed time intervals, e.g. every 2 hours over several consecutive days. Purpose: the conditions of the units in one subgroup are the same, so only common-cause (chance) variation exists within the subgroup. Why n ≤ 5? From the 6th unit onwards the conditions are no longer the same as for the first unit.
3. Why use a control chart? A control chart is used to detect abnormal behavior of the process at an early stage.

## Slide 24
D
I
C
M
A
Control
Define
Z Long term = 6.89
Standard ≥ 4.5
 OK
Z Short term = 7.18 Standard ≥ 4.5
 OK
Zshift = Zst – Zlt = 7.18 – 6.89 = 0.29  Z shift < 1.5
 Proper process control management
Z short term: 7.18 ≥ 4.5 , Z long term: 6.89 ≥ 4.5
 Technology is OK
1.5
Z shift
4.5
C
A
B
D
Lack of management
Lack of skill
Proper management
Lack of skills
Lack of management
Proper skills
Proper management
Proper skills
Control
ZST Or ZST(Bench)
Good
Poor
Technology
4.5 Process Capability review (Z value)
Luminance station
Analyze
Measure
Improve
Notes: Zst: within (short-term), Zlt: overall (long-term). The short-term process capability (Zst) represents the best capability of the current process. Zlt represents the process capability from the customer’s perspective.
Over time the quality deteriorates because of tool and equipment wear, variation between workers and changes of environment. Zshift = Zst − Zlt indicates the degree of deterioration; a shift of about 1.0 to 1.5 is typical. If Zshift is greater than 1.5, management (control) is judged not good; if Zshift is less than 1.5, management is judged good.
Cpk is a measure of the potential capability of the process and equals the minimum of CPU and CPL. Cpk is a ratio that compares two values:
The distance from the process mean to the closest specification limit (USL or LSL)
The one-sided spread of the process (the 3-σ variation) based on the within-subgroup standard deviation
Cpk considers both the location and the variation (within subgroups) of the process.
1. How to select sample for Process Capability? →Select by Rational subgroup
– Lấy theo lot, n con liền nhau của mỗi lot (n<=5), tại các khoảng thời gian cố định. Vd: cách 2 tiếng một lần trong nhiều ngày liên tiếp
Mục đích: để đảm bảo điều kiện
của những sample được lấy là giống hệt nhau, chỉ tồn tại chance cause (nguyên nhân ngẫu nhiên) trong mỗi nhóm.
Tại sao n <= 5? Vì từ con thứ 6 trở đi không đảm bảo điều kiện giống
như con đầu tiên nữa.
How to choose samples for process capability? Take n consecutive units from each lot (n ≤ 5) at fixed time intervals, e.g. every 2 hours over several consecutive days. Purpose: the conditions of the units in one subgroup are the same, so only common-cause (chance) variation exists within the subgroup. Why n ≤ 5? From the 6th unit onwards the conditions are no longer the same as for the first unit.
3. Why use a control chart? A control chart is used to detect abnormal behavior of the process at an early stage.

## Slide 25
Control
Define
4.6 Conclusion
After evaluation, we confirmed that the measurement system at the Luminance station had significant problems during testing. We proved that camera angle misalignment was the root cause of the Luminance NG defect.
We improved the measurement system by correcting the camera angle to 90.0°. We then verified stability and capability using Type 1 Study, Type 3 Gage R&R, Xbar-R control chart, and Process Capability analysis at the Luminance station.
The results demonstrate that the measurement system now achieves high repeatability and no significant bias (Type 1: Cg/Cgk ≥ 1.33; Type 3: %Study Var < 10% and %Contribution < 1%), and the process measured at the Luminance station shows excellent capability (Cpk = 2.39 > 2.0). Result vs goal: Luminance defect rate 226,624 ppm → 0 ppm (0 NG / 420 units after improvement; 95% upper confidence bound ≈ 7,100 ppm, still > 96% below the baseline).
Analyze
Measure
Improve

## Slide 26
5.1 Xbar-R Chart Review after improvement
D
I
C
M
A
Control
Define
Analyze
Measure
Improve
Luminance station
Conclusion:
No out-of-control point and no abnormal pattern (run tests) was found in the Xbar and R control charts. Therefore the process is stable over time.
Date | Sample 1 | Sample 2 | Sample 3 | Sample 4 | Sample 5
1-Jul | 903.99 | 888.65 | 927.84 | 910.28 | 901.09
2-Jul | 910.09 | 898.47 | 930.79 | 906.56 | 901.04
3-Jul | 934.62 | 905.74 | 884.21 | 897.26 | 897.52
4-Jul | 926.69 | 897.88 | 940.66 | 917.92 | 932.89
5-Jul | 901.07 | 919.56 | 916.63 | 920.68 | 896.27
6-Jul | 921.35 | 919.07 | 910.92 | 929.27 | 912.16
7-Jul | 931.37 | 899.74 | 907.23 | 957.93 | 907.04
8-Jul | 902.12 | 906.54 | 893.01 | 893.85 | 907.01
9-Jul | 909.81 | 925.48 | 937.52 | 920.73 | 918.64
10-Jul | 893.27 | 904.75 | 873.01 | 914.26 | 932.42
To monitor the mean and the variation of process after applying improvement and use control chart to monitor process stability over time.
- Sampling method: rational subgroups of the first 5 sets per day at the production line, for 10 days (1-Jul to 10-Jul-2026).
Notes: 1. How to select sample for Process Capability? →Select by Rational subgroup
– Lấy theo lot, n con liền nhau của mỗi lot (n<=5), tại các khoảng thời gian cố định. Vd: cách 2 tiếng một lần trong nhiều ngày liên tiếp
Mục đích: để đảm bảo điều kiện
của những sample được lấy là giống hệt nhau, chỉ tồn tại chance cause (nguyên nhân ngẫu nhiên) trong mỗi nhóm.
Tại sao n <= 5? Vì từ con thứ 6 trở đi không đảm bảo điều kiện giống
như con đầu tiên nữa.
How to choose samples for process capability? Take n consecutive units from each lot (n ≤ 5) at fixed time intervals, e.g. every 2 hours over several consecutive days. Purpose: the conditions of the units in one subgroup are the same, so only common-cause (chance) variation exists within the subgroup. Why n ≤ 5? From the 6th unit onwards the conditions are no longer the same as for the first unit.
3. Why use a control chart? A control chart is used to detect abnormal behavior of the process at an early stage.

## Slide 27
Analyze
Measure
Define
Improve
Control
5.2 Future plan (Management plan)
Management plan
Master Sample Management Procedure: • Inspection Timing: Daily at shift start (08:00~10:00 for day shift, 20:00~22:00 for night shift). • Responsible: PM Technician & Line Leader. • Procedure: 1. Test the Master Sample (reference 910 Cd/m²) at the Luminance station; the reading must be inside the control limits 910 ± 48 Cd/m² (±3s of the Type 1 study after improvement, s = 15.96) and inside spec (800~2000 Cd/m²). 2. If a reading is out of the limits, stop judging OK/NG and report immediately to Line Leader; PE & PM teams will inspect the camera angle (target 90.0°) and confirm with the Master Sample again before restarting. 3. GMES version checking (SW/HW) is skipped during Master Sample testing to focus strictly on hardware measurement stability.
Standardization & horizontal deployment (planned):
• Add the camera-angle check (target 90.0°) to the pre-production setup standard (SOP) of the Luminance station.
• Apply the same Master sample check to the other dual-display models measured with the CS2000 (E3PA Dual, G3 Dual, J1 Dual PPR); the RHD Privacy models also need the 13% luminance compensation (slide 19).
• Keep monitoring with the Xbar-R chart; if a point goes out of control, apply the reaction plan above (step 2).

## Slide 28
Analyze
Measure
Define
Improve
Control
5.3 Result of project
Summary: Luminance defect rate of J1 Dual model before / after improvement
 | Before |  |  |  | After |  | 
Date | 16.Nov.2025 | 2.Jan.2026 | 23.Feb.2026 | 1.Mar.2026 | 15.Jul.2026 | 21.Aug.2026 | 20.Sep.2026
Total NG | 13 | 8 | 92 | 30 | 0 | 0 | 0
Prod Q'ty | 200 | 120 | 240 | 71 | 240 | 120 | 60
Defect rate (ppm) | 65000 | 66666.67 | 383333.33 | 422535.21 | 0 | 0 | 0
Chart: Luminance defect rate of J1 Dual model
| 16.Nov.2025 | 2.Jan.2026 | 23.Feb.2026 | 1.Mar.2026 | 15.Jul.2026 | 21.Aug.2026 | 20.Sep.2026
Total NG | 13 | 8 | 92 | 30 | 0 | 0 | 0
Prod Q'ty | 200 | 120 | 240 | 71 | 240 | 120 | 60
Defect rate (ppm) | 65000 | 66666.6667 | 383333.3333 | 422535.2113 | 0 | 0 | 0
After improvement: 0 NG / 420 sets → 95% upper confidence bound ≈ 7,100 ppm (3/420), still > 96% below the 226,624 ppm baseline. Fisher's exact test, 143/631 before vs 0/420 after: p < 0.001.